## Sound Level Meter 1 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : ACO |  |  | Code | : FI-B-01
Model | : 6236 |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 142016 |  |  | Accuracy | : ± 1.5 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2019-01-09 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM190101 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2019-02-07 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM190201 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2019-03-06 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | SM190301 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2019-04-17 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM190401 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2019-05-09 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | SM190501 | Adjustment @ 94.0,114.0 dB
| 114 | 113.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2019-06-07 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM190601 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-07-13 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200704 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-08-17 | 94 | 93.2 | -0.8 |  | 0.2 | 1 | 1.5 | Pass | Pass | SM200801 | Adjustment @ 94.0,114.0 dB
| 114 | 113.2 | -0.8 |  | 0.2 | 1 | 1.5 | Pass | Pass
2020-10-06 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM201001 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-10-30 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM201003 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-12-18 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | ISTEC201205 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-01-18 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210101 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-03-10 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210302 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-04-15 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210407 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-05-13 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | ISEC210509 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2021-06-15 | 94 | 94.5 | 0.5 |  | 0.2 | 0.7 | 1.5 | Pass | Pass | IICE210610 | Adjustment @ 94.0,114.0 dB
| 114 | 114.5 | 0.5 |  | 0.2 | 0.7 | 1.5 | Pass | Pass
2021-07-22 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210711 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-08-26 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE210812 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-09-24 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210914 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-10-18 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE211017 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-11-23 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | IICE211125 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-12-22 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | EICE211214 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2022-01-24 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220111 | Adjustment @ 94.0,114.2 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-02-16 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | EICE220217 | Adjustment @ 94.0,114.2 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-03-15 | 94 | 94.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | EICE220326 | Adjustment @ 94.0,114.2 dB
| 114 | 114.6 | 0.6 |  | 0.2 | 0.8 | 1.5 | Pass | Pass
2022-04-22 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220436 | Adjustment @ 94.0,114.2 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-05-23 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220544 | Adjustment @ 94.0,114.2 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-06-28 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220654 | Adjustment @ 94.0,114.2 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-07-25 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220759 | Adjustment @ 94.0,114.2 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-08-20 | 94 | 94.5 | 0.5 |  | 0.2 | 0.7 | 1.5 | Pass | Pass | EICE220877 | Adjustment @ 94.0,114.2 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2022-09-19 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | EICE220984 | Adjustment @ 94.0,114.2 dB
| 114 | 113.5 | -0.5 |  | 0.2 | 0.7 | 1.5 | Pass | Pass
2022-10-17 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE221091 | Adjustment @ 94.0,114.2 dB
| 114 | 114.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2022-11-16 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | EICE2211100 | Adjustment @ 94.0,114.2 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-12-16 | 94 | 94.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | EICE2212113 | Adjustment @ 94.0,114.2 dB
| 114 | 114.6 | 0.6 |  | 0.2 | 0.8 | 1.5 | Pass | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.2 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.2 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass

## Sound Level Meter 2 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : ACO |  |  | Code | : FI-B-02
Model | : 6236 |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 142017 |  |  | Accuracy | : ± 1.5 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2019-01-09 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190102 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-02-07 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | SM190202 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2019-03-06 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | SM190302 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2019-04-17 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190402 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-05-09 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190502 | Adjustment @ 94.0,114.0 dB
| 114 | 114.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2019-06-07 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | SM190602 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2019-07-18 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190701 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-08-15 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190801 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-09-16 | 94 | 95.1 | 1.1 |  | 0.2 | 1.3 | 1.5 | Pass | Pass | SM190901 | Adjustment @ 94.0,114.0 dB
| 114 | 115 | 1 |  | 0.2 | 1.2 | 1.5 | Pass | Pass
2019-10-18 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM191001 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-11-16 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | SM191103 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2019-12-13 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM191201 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-01-28 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM200101 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
24/2/2020. | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200203 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-03-20 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | SM200308 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2020-04-20 | 94 | 93.1 | -0.9 |  | 0.2 | 1.1 | 1.5 | Pass | Pass | SM200405 | Adjustment @ 94.0,114.0 dB
| 114 | 113.1 | -0.9 |  | 0.2 | 1.1 | 1.5 | Pass | Pass
2020-05-20 | 94 | 95 | 1 |  | 0.2 | 1.2 | 1.5 | Pass | Pass | SM200501 | Adjustment @ 94.0,114.0 dB
| 114 | 115 | 1 |  | 0.2 | 1.2 | 1.5 | Pass | Pass
2020-06-18 | 94 | 93.3 | -0.7 |  | 0.2 | 0.9 | 1.5 | Pass | Pass | SM200606 | Adjustment @ 94.0,114.0 dB
| 114 | 113.3 | -0.7 |  | 0.2 | 0.9 | 1.5 | Pass | Pass
17/72020 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | SM200701 | Adjustment @ 94.0,114.0 dB
| 114 | 113.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2020-08-17 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200801 | Adjustment @ 94.0,114.0 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-10-06 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM201001 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-10-30 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM201003 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2020-12-18 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISTEC201205 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-01-18 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | ISEC210101 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-03-10 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210302 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-04-15 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | ISEC210407 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-05-13 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | ISEC210509 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-06-15 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210610 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-07-22 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE210711 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-08-26 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | IICE210812 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-09-24 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210914 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-10-18 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | IICE211017 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-11-23 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE211125 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-12-22 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE211214 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-01-22 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220112 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-02-16 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220218 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-03-15 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220327 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-04-22 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220437 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-05-23 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220545 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-06-28 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220655 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-07-25 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220760 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-08-20 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | EICE220878 | Adjustment @ 94.0,114.0 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2022-09-19 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | EICE220985 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2022-10-17 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE221092 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-11-16 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE2211101 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-12-16 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | EICE2212114 | Adjustment @ 94.0,114.0 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass

## Sound Level Meter 3 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : ACO |  |  | Code | : FI-B-03
Model | : 6236 |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 142018 |  |  | Accuracy | : ± 1.5 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2019-01-09 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | SM190103 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2019-02-07 | 94 | 93 | -1 |  | 0.2 | 1.2 | 1.5 | Pass | Pass | SM190203 | Adjustment @ 94.0,114.0 dB
| 114 | 113 | -1 |  | 0.2 | 1.2 | 1.5 | Pass | Pass
2019-03-06 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190303 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-04-17 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190403 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-05-09 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM190503 | Adjustment @ 94.0,114.0 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2019-06-07 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190603 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-07-18 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM190702 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2019-08-15 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM190802 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-09-16 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM190902 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2019-10-18 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM191002 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2019-11-16 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM191104 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2019-12-13 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM191202 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2020-01-28 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200102 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-02-24 | 94 | 94.9 | 0.9 |  | 0.2 | 1.1 | 1.5 | Pass | Pass | SM200204 | Adjustment @ 94.0,114.0 dB
| 114 | 114.9 | 0.9 |  | 0.2 | 1.1 | 1.5 | Pass | Pass
2020-03-20 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200309 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-04-20 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200406 | Adjustment @ 94.0,114.0 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-05-20 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200502 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-06-18 | 94 | 95 | 1 |  | 0.2 | 1.2 | 1.5 | Pass | Pass | SM200607 | Adjustment @ 94.0,114.0 dB
| 114 | 115 | 1 |  | 0.2 | 1.2 | 1.5 | Pass | Pass
2020-07-17 | 94 | 93.5 | -0.5 |  | 0.2 | 0.7 | 1.5 | Pass | Pass | SM200702 | Adjustment @ 94.0,114.0 dB
| 114 | 113.5 | -0.5 |  | 0.2 | 0.7 | 1.5 | Pass | Pass
2020-08-17 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200801 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-10-06 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM201001 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2020-10-30 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM201003 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2020-12-18 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISTEC201205 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-01-18 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210101 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-03-10 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210302 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-04-15 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210407 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-05-13 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210509 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-06-15 | 94 | 94.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | IICE210610 | Adjustment @ 94.0,114.0 dB
| 114 | 114.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2021-07-22 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | IICE210711 | Adjustment @ 94.0,114.0 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2021-08-26 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210812 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-09-24 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE210914 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-10-18 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE211017 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-11-23 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE211125 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-12-22 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE211214 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-01-22 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220113 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-02-16 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220219 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-03-15 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | EICE220328 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2022-04-22 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220438 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-05-23 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220546 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-07-25 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220761 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-08-20 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220879 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-09-19 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | EICE220986 | Adjustment @ 94.0,114.0 dB
| 114 | 113.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2022-10-17 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE221093 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-11-16 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | EICE2211102 | Adjustment @ 94.0,114.0 dB
| 114 | 114.3 | 0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2022-12-16 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE2212115 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass

## Form
Measuring Equipment Record
Name | [5 empty] | Manufacturer |  |  |  |  | Code
Model | [5 empty] | Range/Resol |  |  |  |  | Calibration interval
Serial no. | [5 empty] | Accuracy |  |  |  |  | Accessory
Date Cal. | Nominal value | Reding value | Error (1) | Uncert. (2) | (1) + (2) | MPE | Judgement | Cert. no / Remark | Verify By | Repairing Record
[10 empty] | Date | Detail | Company Name | Cost

## Sound Level Meter 4 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : ACO |  |  | Code | : FI-B-12
Model | : 6226 |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 200006 |  |  | Accuracy | : ± 1.5 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2020-02-24 | 94 | 93.3 | -0.7 |  | 0.2 | 0.9 | 1.5 | Pass | Pass | SM200205 | Adjustment @ 94.0,114.0 dB
| 114 | 113.3 | -0.7 |  | 0.2 | 0.9 | 1.5 | Pass | Pass
2020-03-20 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200310 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-04-20 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200407 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-05-20 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200503 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2020-06-18 | 94 | 95.4 | 1.4 |  | 0.2 | 1.6 | 1.5 | Fail | Pass | SM200608 | Adjustment @ 94.0,114.0 dB
| 114 | 115.4 | 1.4 |  | 0.2 | 1.6 | 1.5 | Fail | Pass
2020-07-17 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200703 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-08-17 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200801 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-10-06 | 94 | 94.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | SM201001 | Adjustment @ 94.0,114.0 dB
| 114 | 114.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2020-10-30 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM201003 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2020-12-18 | 94 | 93.2 | -0.8 |  | 0.2 | 1 | 1.5 | Pass | Pass | ISTEC201205 | Adjustment @ 94.0,114.0 dB
| 114 | 113.2 | -0.8 |  | 0.2 | 1 | 1.5 | Pass | Pass
2021-01-18 | 94 | 94.7 | 0.7 |  | 0.2 | 0.9 | 1.5 | Pass | Pass | ISEC210101 | Adjustment @ 94.0,114.0 dB
| 114 | 114.7 | 0.7 |  | 0.2 | 0.9 | 1.5 | Pass | Pass
2021-03-10 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210302 | Adjustment @ 94.0,114.0 dB
| 114 | 114.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2021-04-15 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210407 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-05-13 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210509 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-06-15 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | IICE210610 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-07-22 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE210711 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-08-26 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210812 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-09-24 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | IICE210914 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2021-10-18 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE211017 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-11-23 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | IICE211125 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-12-22 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE211214 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-01-24 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220114 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-02-16 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | EICE220220 | Adjustment @ 94.0,114.0 dB
| 114 | 113.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2022-03-15 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220329 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-04-22 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220439 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-05-23 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220547 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-06-28 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220656 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-07-25 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220762 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-08-20 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220880 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-09-19 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220987 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-10-17 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | EICE221094 | Adjustment @ 94.0,114.0 dB
| 114 | 113.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2022-11-16 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE2211103 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-12-16 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE2212116 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass

## Sound Level Meter 5 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : ACO |  |  | Code | : FI-B-13
Model | : 6226 |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 200055 |  |  | Accuracy | : ± 1.5 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2020-08-17 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200801 | Adjustment @ 94.0,114.0 dB
| 114 | 113.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass
2020-10-06 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass | SM201001 | Adjustment @ 94.0,114.0 dB
| 114 | 113.6 | -0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2020-10-30 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM201003 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2020-12-18 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISTEC201205 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-01-18 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210101 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-03-10 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | ISEC210302 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-04-15 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210407 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-05-13 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210509 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2021-06-15 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE216010 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-07-22 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210711 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-08-26 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210812 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-09-24 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | IICE210914 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2021-10-18 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE211017 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2021-11-23 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE211125 | Adjustment @ 94.0,114.0 dB
| 114 | 114.4 | 0.4 |  | 0.2 | 0.6 | 1.5 | Pass | Pass
2021-12-22 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE211214 | Adjustment @ 94.0,114.0 dB
| 114 | 113.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-01-22 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220115 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-02-16 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220221 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-03-15 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220330 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-04-22 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220440 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-05-23 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220548 | Adjustment @ 94.0,114.0 dB
| 114 | 114.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass
2022-06-28 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220657 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-07-25 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220763 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-08-20 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220881 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
2022-09-19 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220988 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-10-17 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE221095 | Adjustment @ 94.0,114.0 dB
| 114 | 113.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-11-16 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE2211104 | Adjustment @ 94.0,114.0 dB
| 114 | 114.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass
2022-12-16 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE2212117 | Adjustment @ 94.0,114.0 dB
| 114 | 114 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0,114.0 dB
| 114 |  | -114 |  | 0.2 | 114.2 | 1.5 | Fail | Pass

## Sound Level Meter 6 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : Scarlet Tech |  |  | Code | : FI-B-30
Model | : ST-21D |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 820720 |  |  | Accuracy | : ± 0.7 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2022-09-19 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 0.7 | Pass |  | EICE220989 | Adjustment @ 94.0 dB
2022-10-17 | 94 | 94 | 0 |  | 0.2 | 0.2 | 0.7 | Pass |  | EICE221096 | Adjustment @ 94.0 dB
2022-11-16 | 94 | 94.3 | 0.3 |  | 0.2 | 0.5 | 0.7 | Pass | Pass | EICE2211105 | Adjustment @ 94.0 dB
2022-12-16 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 0.7 | Pass | Pass | EICE2212118 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB

## Sound Level Meter 7 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : Scarlet Tech |  |  | Code | : FI-B-33
Model | : ST-21D |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 820745 |  |  | Accuracy | : ± 0.7 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2022-12-16 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 0.7 | Pass | Pass | EICE2212119 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB

## Sound Level Meter 8 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : Scarlet Tech |  |  | Code | : FI-B-34
Model | : ST-21D |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 820746 |  |  | Accuracy | : ± 0.7 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2022-12-16 | 94 | 93.6 | -0.4 |  | 0.2 | 0.6 | 0.7 | Pass | Pass | EICE2212120 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB

## Sound Level Meter 9 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : Scarlet Tech |  |  | Code | : FI-B-35
Model | : ST-21D |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 820747 |  |  | Accuracy | : ± 0.7 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2022-12-16 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 0.7 | Pass | Pass | EICE2212121 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB

## Sound Level Meter 10 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : Scarlet Tech |  |  | Code | : FI-B-36
Model | : ST-21D |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 820748 |  |  | Accuracy | : ± 0.7 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2022-12-16 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 0.7 | Pass | Pass | EICE2212122 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB

## Sound Level Meter 11 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : Scarlet Tech |  |  | Code | : FI-B-37
Model | : ST-21D |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 820749 |  |  | Accuracy | : ± 0.7 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2022-12-16 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 0.7 | Pass | Pass | EICE2212123 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB

## Sound Level Meter 12 (hidden)
Measuring Equipment Record
Name | : Sound Level Meter |  |  | Manufacturer | : Scarlet Tech |  |  | Code | : FI-B-38
Model | : ST-21D |  |  | Range / Resol. | : 28 - 130 dB |  |  | Calibration interval | : 1 Month
Serial no. | : 820750 |  |  | Accuracy | : ± 0.7 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2022-12-16 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 0.7 | Pass | Pass | EICE2212124 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 0.7 | Fail | Pass |  | Adjustment @ 94.0 dB

## Noise Dose Meter 0 (hidden)
Measuring Equipment Record
Name | : Noise Dose Meter |  |  | Manufacturer | : Tenmars |  |  | Code | : FI-B-14
Model | : ST-130 |  |  | Range / Resol. | : 30 - 130 dB |  |  | Calibration interval | : 1 Year
Serial no. | : 190500027 |  |  | Accuracy | : ± 1.5 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2020-04-21 | 94 | 91.9 | -2.1 |  | 0.2 | 2.3 | 1.5 | Fail |  | SM200415 | Adjustment @ 94.0 dB
2020-06-12 | 94 | 93.2 | -0.8 |  | 0.2 | 1 | 1.5 | Pass |  | SM200604 | Adjustment @ 94.0 dB
2020-07-13 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM200704 | Adjustment @ 94.0 dB
2020-08-18 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM200802 | Adjustment @ 94.0 dB
2020-10-16 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM201002 | Adjustment @ 94.0 dB
2020-12-18 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | ISTEC201205 | Adjustment @ 94.0 dB
2021-04-15 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210408 | Adjustment @ 94.0 dB
2021-05-13 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210509 | Adjustment @ 94.0 dB
2021-09-24 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | IICE210914 | Adjustment @ 94.0 dB
2021-11-13 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC211123. | Adjustment @ 94.0 dB
2022-01-14 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220105 | Adjustment @ 94.0 dB
2022-04-27 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | EICE220441 | Adjustment @ 94.0 dB
2022-07-27 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220768 | Adjustment @ 94.0 dB
2022-10-17 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE221097 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB

## Noise Dose Meter 00 (hidden)
Measuring Equipment Record
Name | : Noise Dose Meter |  |  | Manufacturer | : Tenmars |  |  | Code | : FI-B-15
Model | : ST-130 |  |  | Range / Resol. | : 30 - 130 dB |  |  | Calibration interval | : 3 Month
Serial no. | : 190500064 |  |  | Accuracy | : ± 1.5 dB |  |  | Accessory | : Battery, Windscreen, Microphone
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no | Remark
2020-04-21 | 94 | 91.5 | -2.5 |  | 0.2 | 2.7 | 1.5 | Fail |  | SM200414 | Adjustment @ 94.0 dB
2020-06-12 | 94 | 93.1 | -0.9 |  | 0.2 | 1.1 | 1.5 | Pass |  | SM200605 | Adjustment @ 94.0 dB
2020-07-13 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | SM200704 | Adjustment @ 94.0 dB
2020-08-18 | 94 | 93.7 | -0.3 |  | 0.2 | 0.5 | 1.5 | Pass | Pass | SM200802 | Adjustment @ 94.0 dB
2020-10-16 | 94 | 93.8 | -0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | SM201002 | Adjustment @ 94.0 dB
2020-12-18 | 94 | 93.5 | -0.5 |  | 0.2 | 0.7 | 1.5 | Pass | Pass | ISTEC201205 | Adjustment @ 94.0 dB
2021-04-15 | 94 | 93.9 | -0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC210408 | Adjustment @ 94.0 dB
2021-05-13 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | ISEC210509 | Adjustment @ 94.0 dB
2021-09-24 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | IICE210914 | Adjustment @ 94.0 dB
2021-11-13 | 94 | 94.1 | 0.1 |  | 0.2 | 0.3 | 1.5 | Pass | Pass | ISEC211123 | Adjustment @ 94.0 dB
2022-01-14 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220106 | Adjustment @ 94.0 dB
2022-04-27 | 94 | 94.2 | 0.2 |  | 0.2 | 0.4 | 1.5 | Pass | Pass | EICE220442 | Adjustment @ 94.0 dB
2022-07-27 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE220769 | Adjustment @ 94.0 dB
2022-10-17 | 94 | 94 | 0 |  | 0.2 | 0.2 | 1.5 | Pass | Pass | EICE221098 | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB
| 94 |  | -94 |  | 0.2 | 94.2 | 1.5 | Fail | Pass |  | Adjustment @ 94.0 dB

## Sound Level Calibrator (hidden)
Measuring Equipment Record
Name | :Sound Level Calibrator | [5 empty] | Manufacturer | : Tenmars |  |  |  | Code | :
Model | TM-100 | [5 empty] | Range / Resol. | : 94.0 และ 114.0 dB |  |  |  | Calibration interval | : 1 Year
Serial no. | : 160100571 | [5 empty] | Accuracy | : ± 0.5 dB |  |  |  | Accessory
Date | Norminal value | Reding value | Error (1) | Uncert. (2) | (1) + (2) | MPE | Judgement |  | Cert. no | Remark | Repairing Record
[11 empty] | Date | Detail | Company Name | Cost
2017-08-24 | 94 | 93.9 | -0.1 |  | 0.1 | 0.5 | Pass | Pass
| 114 | 114 | 0 |  | 0 | 0.5 | Pass | Pass
2018-09-05 | 94 | 94.1 | 0.1 | 0.1 | 0.2 | 0.5 | Pass | Pass | 0376SV18
| 114 | 114.04 | 0.04 | 0.1 | 0.14 | 0.5 | Pass | Pass
2020-09-09 | 94 | 94.09 | 0.09 | 0.1 | 0.19 | 0.5 | Pass | Pass | 0325SV20
| 114 | 114.05 | 0.05 | 0.1 | 0.15 | 0.5 | Pass | Pass
2021-12-02 | 94 | 94.2 | 0.2 | 0.1 | 0.3 | 0.5 | Pass | Pass | CP20210080EB
| 114 | 114.13 | 0.13 | 0.1 | 0.23 | 0.5 | Pass | Pass
2022-12-06 | 94 | 94.16 | 0.16 | 0.1 | 0.26 | 0.5 | Pass | Pass | CP20220359EA
| 114 | 114.13 | 0.13 | 0.1 | 0.23 | 0.5 | Pass | Pass

## pH Meter 1 (hidden)
Measuring Equipment Record
Name | :pH Meter |  |  | Manufacturer | Extech |  |  | Code | :
Model | : PH-100 |  |  | Range / Resol. | : pH 0 - 14 |  |  | Calibration interval | : 1 Month
Serial no. | : 398397 |  |  | Accuracy | ± 0.01 pH |  |  | Accessory | : Battery, Cover
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no / Remark
2018-11-24 | 4 | 4.01 | 0.01 |  | 0 | 0.01 | 0.01 | Pass | Pass
| 7 | 6.92 | -0.08 |  | 0 | 0.08 | 0.01 | False | Pass
| 10 | 10.08 | 0.08 |  | 0 | 0.08 | 0.01 | False | Pass
241821 | 4
7
10
241873 | 4
7
10
241904 | 4
7
10
241921 | 4
7
10
241970 | 4
7
10
241993 | 4
7
10
242024 | 4
7
10
หมายเลขแบบฟอร์ม : PM-LA-500-04 / วันที่ใช้งาน : 14-07-2018 / แก้ไขครั้งที่ : 00

## pH Meter 2 (hidden)
Measuring Equipment Record
Name | :pH Meter |  |  | Manufacturer | ADWA |  |  | Code | :
Model | : AD-11 |  |  | Range / Resol. | : pH 0 - 14 |  |  | Calibration interval | : 1 Month
Serial no. |  |  |  | Accuracy | : '±0.1 pH ±0.5˚C (25˚C) |  |  | Accessory | : Battery, Cover
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no / Remark
241755 | 4 | 4.1 | 0.1 |  | 0 |  | 0.1 | Pass | Pass
| 7 | 7.1 | 0.1 |  | 0 |  | 0.1 | Pass
| 10 | 10.1 | 0.1 |  | 0 |  | 0.1 | Pass
ชำรุด ในเดือน ธันวาคม 2561
หมายเลขแบบฟอร์ม : PM-LA-500-04 / วันที่ใช้งาน : 14-07-2018 / แก้ไขครั้งที่ : 00

## pH Meter 3 (hidden)
Measuring Equipment Record
Name | :pH Meter |  |  | Manufacturer | Smart Sensor |  |  | Code | :
Model | : AS218 |  |  | Range / Resol. | : 0.00 - 14.00 pH |  |  | Calibration interval | : 1 Month
Serial no. | : 3806615 |  |  | Accuracy | : ±0.03 pH |  |  | Accessory | : Battery,,
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no / Remark
241867 | 4 | 4.01 | 0.01 |  | 0 | 0.01 | 0.03 | Pass | Pass
| 6.86 | 6.84 | -0.02 |  | 0 | 0.02 | 0.03 | Pass
| 9.18 | 9.19 | 0.01 |  | 0 | 0.01 | 0.03 | Pass
242024 | 4
7
10
หมายเลขแบบฟอร์ม : PM-LA-500-04 / วันที่ใช้งาน : 14-07-2018 / แก้ไขครั้งที่ : 00

## DO Meter 1 (hidden)
Measuring Equipment Record
Name | : DO Meter |  |  | Manufacturer | Lutron |  |  | Code | :
Model | : PDO-519 |  |  | Range / Resol. | : 0 - 20 mg/L 0 - 100% O2 |  |  | Calibration interval | : 1 Month
Serial no. | : 92077 |  |  | Accuracy | : ' ±0.4 mg/L ± 0.7 % O2 |  |  | Accessory | : Battery, Cover
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no / Remark
241755 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass |  | เติมน้ำยา
241762 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241795 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241824 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241852 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241883 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241914 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241944 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241974 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
242005 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
242036 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
242066 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass |  | เติมน้ำยา 16/10/2562
หมายเลขแบบฟอร์ม : PM-LA-500-04 / วันที่ใช้งาน : 14-07-2018 / แก้ไขครั้งที่ : 00

## DO Meter 2 (hidden)
Measuring Equipment Record
Name | : DOMeter |  |  | Manufacturer | Lutron |  |  | Code | :
Model | : PDO-519 |  |  | Range / Resol. | : 0 - 20 mg/L 0 - 100% O2 |  |  | Calibration interval | : 1 Month
Serial no. | : 92078 |  |  | Accuracy | : ±0.4 mg/L ± 0.7 % O2 |  |  | Accessory | : Battery, Cover
Date | Norminal value | Reding value | Error (1) |  | Uncert. (2) | (1) + (2) | Accuracy | Judgement |  | Cert. no / Remark
241755 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass |  | เติมน้ำยา
241762 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241795 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241824 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241852 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241883 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241914 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241944 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
241974 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
242005 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
242036 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass
242066 | 20.9 | 20.9 | 0 |  | 0 | 0 | 0.7 | Pass |  | เติมน้ำยา 16/10/2562
หมายเลขแบบฟอร์ม : PM-LA-500-04 / วันที่ใช้งาน : 14-07-2018 / แก้ไขครั้งที่ : 00